## Slide 1: COMPARATIVE STUDY OF TWO TECHNIQUES TO MEASURE HIV-1 VIRAL LOAD IN A FRENCH COMMERCIAL LABORATORY G. KREPLAK, R. LEGOUGE Centre biologique du Chemin Vert, Paris, France
P676
ABSTRACT
Objectives: The aim of the study was to compare the performance of the NucliSens EasyQ HIV-1 V1.1 assay (bioMérieux, Boxtel, The Netherlands), a real time NASBA amplification assay for the quantitative determination of HIV-1 RNA in clinical specimen with the Versant HIV-1 RNA 3.0 (bDNA, Bayer Diagnostics) technique routinely used in our laboratory.
Methods: A total of 221 plasma samples were collected from HIV positive / AIDS patients. For each sample, HIV-1 viral load was measured using the NucliSens EasyQ HIV-1 V1.1 assay and the assay we use for routine i.e. Versant HIV-1 RNA 3.
An 8-member calibration panel from Versant was used to check for proper set-up and functioning of both systems.
Results: Concordant results were obtained in 202/221 samples (91.40%).
Seventeen results were found to be discordant. Among them, 6 were negative with Versant HIV-1 RNA 3 and quantified with EasyQ HIV-1 V1.1 and 11 were negative with EasyQ HIV-1 V1.1 giving a quantification result with Versant HIV-1 RNA 3.
These discordant results showed very low values of viral load, always inferior to 2.89 log, corresponding to < 776 copies/ml.
Two EasyQ HIV-1 V1.1 results were considered as invalid.
The results obtained demonstrate a good correlation between the two techniques for both low and high viral loads.
Conclusion: Analysis of a set of 221 clinical routine samples revealed that EasyQ HIV-1 V1.1 has a comparable performance to the Versant HIV-1 RNA 3 assay used in our laboratory. No difference in quantitation was observed between NucliSens EasyQ HIV-1 and Versant HIV-1 RNA 3.
These performance results would enable us to switch from our routine method to NucliSens EasyQ HIV-1 without any impact on patient follow-up.
OBJECTIVES
Compare the routine laboratory technique (Versant HIV bdna) with a real-time PCR technique (Nuclisens EasyQ HIV-1) on unselected laboratory samples of patient plasma.
Attempt to evaluate the possibility of replacing our current technique with a real-time technique.
EQUIPMENT AND METHODS
1. Samples
220 samples of plasma collected by our laboratory were analyzed using both techniques in several different sets of analyses.
Versant: 3 sets of analyses of 84 samples.
Nuclisens: 6 sets with samples of different sizes.
2. Controls
Versant: 1 set of calibration solutions and 3 controls of various levels for each set.
Nuclisens: 1 extraction, amplification and calibration control per sample with a calculation of the respective quantities of patient RNA and control RNA using two mathematical algorithms managed by the software of the EASYQ system.
3. Nucleic acid extraction
Versant: ultracentrifugation for one hour to concentrate the volume of HIV-RNA.
Nuclisens: extractor automated extraction system (bioMérieux).
4. Detection
Versant: signal amplification
Nuclisens: real-time target amplification and detection by monitoring the target and control kinetics on the EASYQ system.
220 plasma samples were tested using both techniques.
Among them,
- 107 were positive with both techniques
- 94 were negative with both techniques
Giving an overall similarity of 91.36%.
Seventeen samples do not match:
- Six are negative using the Versant technique, but assayed using the Nuclisens technique.
- Eleven are negative using the Nuclisens technique, but assayed using the Versant technique.
Each sample has a low viral load of always less than 2.89 log.
We hypothesize that these differences may be related to a detection limit for certain subtypes depending on the technique used.
Two samples are invalidated by the Nuclisens technique.
The overly low amplification of the control for 1 sample may be explained by the presence of inhibitors. One is probably linked to the absence of any control being added.
Two samples whose values were greater than the versant assay technique zone were diluted in negative sera and re-assayed. The results thus obtained are fully comparable to those obtained directly using the Nuclisens technique.
CONCLUSION
The performance of both techniques are fully comparable. The Nuclisens technique is extremely fast. Nevertheless, we have retained the routine technique we currently use due to the fact that the extraction technique is not suited to our daily throughput. Knowing that bioMérieux has launched a new semi-automated extraction system and is developing a fully automated addressing small, medium and large volumes., further studies should be performed.
RESULTS AND DISCUSSION

## Slide 2: Objectif de l’étude
Comparaison de 2 techniques de charge virale VIH:
bioMérieux
NucliSens EasyQ HIV-1 temps réel
Bayer Diagnostics
Versant HIV-1 RNA

## Slide 3: Méthode
Étude de 221 échantillons de patients
3 séries sur la technique de routine Versant
6 séries sur la plate-forme EasyQ

## Slide 4: Résultats
107 positifs avec les 2 techniques
94 négatifs avec les 2 techniques
17 résultats discordants ( < 776 copies / ml)
Correlation : 91,36 %

## Slide 5: Courbe de corrélation

## Slide 6: Conclusion
Les performances des 2 techniques sont entièrement comparables
La technique NucliSens temps réel a été trouvée extrêmement rapide et agréable d’utilisation
L’automate Extractor n’est par contre pas adapté au volume du laboratoire (12000 tests par an)

## Slide 7: Matériels NucliSens